ABSTRACT

Objective This study aimed to evaluate the accuracy in terms of trueness and precision of eight IOS and the effect of different finishing line designs on the IOS’s accuracy.

Materials and Methods Three printed models of the maxillary arch with maxillary right first molar virtually prepared with chamfer, shoulder, and vertical preparation designs were used as master models in this study. Each model was scanned thirty times with each IOS: Medit®i700, Planscan Emerald®S, CEREC®Primescan, TRIOS®3, CS3600®, MEDIT®i500, Heron™ 3Disc, and Cerec®Omnicam. The trueness was measured by superimposition of the scanned dataset made with IOS and the scanned dataset made with a lab scanner (In Lab Medit T710) that was used as a reference and the deviation was measured and expressed as a color-coded map by the metrology program (Medit compare, version 2.3.5.892.), while precision was measured by the superimposition of the scans of each IOS on each other.

Statistical Analysis The data were analyzed statistically using repeated measure ANOVA test, one-way ANOVA test, Bonferroni test, and Tukey’s test at significance level of 0.05.

Results The tested IOS showed significant differences in trueness and precision. Medit®i700and CEREC®Primescan recorded the highest precision with no significant difference between them, while Medit®i700recorded the highest trueness as compared to other IOS s. Each IOS showed significant differences in trueness and precision with the three finishing line designs except CEREC®Primescan and Heron™ 3Disc which showed no significant difference in trueness with the three finishing line designs and CS3600® which showed no significant difference in precision with the three finishing line designs.

Conclusion A significant difference in accuracy was found among the tested IOSs and the type of finishing line design had a significant effect on IOS’s accuracy.